REFILL INSTRUCTIONS FOR ZAG CANISTERS

DISCONNECT ELECTRICAL CORD, OUTPUT, AND DRAIN.  ALLOW SYSTEM TO COOL.

A198 – ZEOLITE (MB105)	Media color is white

A199 – PURAFIL / PURAKOL (MB106)	Media color is black

A199 – PURAFIL / SELECT (MB107)  Media color is purple

Remove cover on system

Place system on side and remove 4 nuts that mount each sub–assembly

Note:	Assemblies should not fall out, tubing will hold in place

Place system upright and remove hardware from side panels

A198 Canister Removal:

Fittings are push collar to remove and install tubing

Disconnect tubing at tee and remove the input tee from top of canister

Raise assembly to disconnect tubing from valve on back of bracket

Remove canisters from bracket

Replace media as outlined on page 2

Re-install canisters, and reconnect as originally configured

A199 Canister Removal:

Remove tubing on top of canisters using a 9/16 wrench

Remove MB107 canister, pull out of bracket, and disconnect tubing from bottom of canister

Raise bracket to remove MB106 canister from bracket and disconnect tubing from bottom of canister

Replace media as outlined on page 2

Ensure MB107 canister is on outside of system, flow arrows are up, and tubing connects to coil

Ensure MB106 flow arrows are down and tubing connects to regulator on front panel

Re-install canisters, and reconnect as originally configured

Secure brackets to chassis

Secure side panels and cover

RUN SYSTEM TO PURGE FOR 2 HOURS BEFORE USE

REPLACING MEDIA AND FILTERS IN ALL CANISTERS

Each canister needs to be re-built independently.  Caps must be re-installed on original side of canister.

Removing filters, o – rings, and media:

Unscrew caps with 7/8 wrench

Carefully remove felt filter

Dispose of media, felt filters, and o – rings

Replacement of filters, o-rings, and media:

Thoroughly clean inside of canister and threads on the caps

Use the pre-lubricated replacement 0-rings and place on caps

Thread one cap on canister and tighten

Install replacement felt filter into canister flat against capped end

Replace with correct media as labeled on canister

Fill canister 3/4 inch from top

Press felt filter flat against the media and tighten top cap

Note:  Ensure media is not outside of felt filter before capping canister, or it will contaminate and or clog system output